Sample Requirements for Biocompatibility Testing (per extract)

| Test Method | *Minimum Sample Requirement for: | **Turnaround Time |
| --- | --- | --- |
| Cytotoxicity | material thickness of ≥0.5 mm = 60 cm2 material thickness of <0.5 mm = 120 cm2 4 grams if surface area cannot be determined | 2 - 3 weeks |
| GPMT (Kligman) Sensitization | material thickness of ≥0.5 mm = 180 cm2 material thickness of <0.5 mm = 360 cm2 12 grams if surface area cannot be determined | 7 - 9 weeks |
| Closed Patch (Buehler) | material thickness of ≥0.5 mm = 360 cm2 material thickness of <0.5 mm = 720 cm2 24 grams if surface area cannot be determined | 7 – 9 weeks |
| Primary Skin Irritation | material thickness of ≥0.5 mm = 60 cm2 material thickness of <0.5 mm = 120 cm2 4 grams if surface area cannot be determined | 2 – 3 weeks |
| Intracutaneous Reactivity | material thickness of ≥0.5 mm = 60 cm2 material thickness of <0.5 mm = 120 cm2 4 grams if surface area cannot be determined | 2 – 3 weeks |
| Systemic Toxicity; acute | material thickness of ≥0.5 mm = 60 cm2 material thickness of <0.5 mm = 120 cm2 4 grams if surface area cannot be determined | 2 – 3 weeks |
| Implantation | 18 pcs. (1 – 3 mm) x 10 mm | Various time frames |
| Hemocompatibility; Hemolysis | 270 cm2 ≥0.5 mm thick 540 cm2 <0.5 mm thick and at least 3 units per method requested | 2 – 3 weeks |
| Material Characterization |  |  |
| Infrared analysis | 10 grams | 3 – 5 days |
| Physicochemical Test | 600 cm2 or 20 grams | 1 – 2 weeks |

*Please contact our customer service department if sample requirements are not attainable.  If possible, additional material is desired to avoid delay in processing due to possible sample shortage.

**Turn around time is based on actual set-up and test time.  Please call prior to sample submittal to confirm turn around times.