Evaluation Tests to Consider - ISO 10993-1:2003*

Section 6: Selection of biological evaluation tests; Tables 1 / 2

| Device Categorized as: |  |  | Biological Effects |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Body Contact |  | Contact Duration Limited •Less than 24 hours Prolonged •24 hours to 30 days Permanent •Over 30 days | Cytotoxicity | Sensitization | Irritation or Intracutaneous Reactivity | Systemic Toxicity (acute) | Subacute and Subchronic Toxicity | Genotoxicity | Implantation | Hemocompatibility | Chronic Toxicity | Carcinogenicity |
| Surface Device | Skin | Limited | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Permanent | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  | Mucosal Membrane | Limited | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Permanent | ♦ | ♦ | ♦ |  | ♦ | ♦ |  |  |  |  |
|  | Breached or compromised surfaces | Limited | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Permanent | ♦ | ♦ | ♦ |  | ♦ | ♦ |  |  |  |  |
| External communicating device | Blood Path, Indirect | Limited | ♦ | ♦ | ♦ | ♦ |  |  |  | ♦ |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ | ♦ |  |  |  | ♦ |  |  |
|  |  | Permanent | ♦ | ♦ |  | ♦ | ♦ | ♦ |  | ♦ | ♦ | ♦ |
|  | Tissue/Bone/ Dentin | Limited | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |  |  |  |
|  |  | Permanent | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |  | ♦ | ♦ |
|  | Circulating Blood | Limited | ♦ | ♦ | ♦ | ♦ |  |  |  | ♦ |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |  |  |
|  |  | Permanent | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |
| Implant device | Tissue/Bone | Limited | ♦ | ♦ | ♦ |  |  |  |  |  |  |  |
|  |  | Prolonged | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |  |  |  |
|  |  | Permanent | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |  |  |  |
|  | Blood | Limited | ♦ | ♦ | ♦ | ♦ | ♦ |  | ♦ | ♦ | ♦ |  |
|  |  | Prolonged | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |  |
|  |  | Permanent | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ | ♦ |

*Additional tests may be required to satisfy FDA requirements.